Release Notes

ProFlex™ PCR System Firmware v2.2.2

For the Applied Biosystems® ProFlex™ PCR System

Contents

OVERVIEW

FEATURES

DOWNLOAD INSTRUCTIONS

UPGRADE INSTRUCTIONS

LEGAL STATEMENTS

Overview

Firmware v2.2.2 for the ProFlex™ PCR System

This firmware upgrade is for ProFlex™ PCR Systems only:

4483636	ProFlex™ Base Module

4483637	ProFlex™ 96-Well Sample Block
4483638	ProFlex™ 3x32-Well Sample Block
4484071	ProFlex™ 2x96-Well Sample Block
4484074	ProFlex™ 2xFlat Sample Block
4484072	ProFlex™ 2x384 Sample block

4484075	ProFlex™ 96-Well PCR System
4484073	ProFlex™ 3x32-Well PCR System
4484076	ProFlex™ 2x96-Well PCR System
4484078	ProFlex™ 2xFlat PCR System
4484077	ProFlex™ 2x384-Well PCR System

Firmware version 2.2.2 will include all enhancements of previous releases in addition to the new features listed below on the new hardware platform.  This version will not in any way affect the performance of your ProFlex™ system except for the specific changes outlined below.

New Features/Changes of version 2.2.2

- Touchscreen sensitivity tuning for improved response

Note: If you have cloud accounts linked up already in firmware version 2.x.x, please perform administrator cloud account unlink from the cloud before you perform the firmware update. You will required to perform the cloud link up again when upgraded to this new version.

This firmware version 2.2.2 is not compatible with Thermal Cycler Fleet Control Software version 1.1.0 and below. If you are using this software for ProFlex instruments control, DO NOT upgrade the instrument to this version. There are plans for new fleet control software release to support this firmware version.

This firmware version 2.2.2 is not compatible with ProFlex Firmware version 1.X.X. If you are using firmware 1.X.X, DO NOT upgrade the instrument to this version.

Downloading the firmware file

Download the attached file onto a USB memory stick (shown in .update format).

- Important: For Internet Explorer users, right click the above “download” link.  Click “Save Target As”.  Change the file extension to “.update” and click save

- Save/Transfer the .update file (not in any folder) on your USB stick.

- Make sure that no other .update files are present on the memory stick.

Updating your ProFlex™ instrument

- Power on the instrument.

- Insert the USB memory stick containing the firmware update into one of the USB ports in the ProFlex™ instrument.

- On the Home screen, touch ‘Settings’.

- In the Settings menu, touch ‘Maintenance & Services’.

- Touch 'Software Updates’.

- If you have downloaded the firmware properly, the file will appear on the screen.

- Select the file on the screen, and press 'Upgrade' at the bottom of the screen.

- Touch ‘Yes’ to confirm that you want to upgrade.

IMPORTANT! Ensure that you do not turn off the instrument when the upgrade is in progress.

- When the upgrade process is complete, the instrument will restart automatically. When it has rebooted, you can remove your USB stick.

LEGAL STATEMENTS

Important licensing information:

These products may be covered by one or more Limited Use Label Licenses. By use of these products, you accept the terms and conditions of all applicable Limited Use Label Licenses.

The trademarks mentioned herein are the property of Life Technologies Corporation and/or its affiliate(s) or their respective owners.

LIFE TECHNOLOGIES CORPORATION AND/OR ITS AFFILIATE(S) DISCLAIM ALL WARRANTIES WITH RESPECT TO THIS DOCUMENT, EXPRESSED OR IMPLIED, INCLUDING BUT NOT LIMITED TO THOSE OF MERCHANTABILITY, FITNESS FOR A PARTICULAR PURPOSE, OR NON-INFRINGEMENT. TO THE EXTENT ALLOWED BY LAW, IN NO EVENT SHALL LIFE TECHNOLOGIES AND/OR ITS AFFILIATE(S) BE LIABLE, WHETHER IN CONTRACT, TORT, WARRANTY, OR UNDER ANY STATUTE OR ON ANY OTHER BASIS FOR SPECIAL, INCIDENTAL, INDIRECT, PUNITIVE, MULTIPLE OR CONSEQUENTIAL DAMAGES IN CONNECTION WITH OR ARISING FROM THIS DOCUMENT, INCLUDING BUT NOT LIMITED TO THE USE THEREOF.

For Research Use Only. Not for use in diagnostic procedures.

© Copyright 2018 Life Technologies Corporation. All rights reserved.